EEG Safety Procedure

Faculty of Psychology and Neuroscience Maastricht University

Franc Donkers, Fren Smulders, Conny Quaedflieg, Lars Hausfeld, Michael Schwartze 	Maastricht, 09-09-2025

------------------------------------------------------------------------------------------------------------------------

In this document the following is discussed:

- Aim

- Description of the EEG technique

- Known persistent side-or after-effects

- Starting an EEG study

- Experimental procedures

- Safety guidelines and procedures

  - Electrical Safety

  - Reducing Risk for Disease Transmission

  - Chemical safety

  - Specific instructions for disinfection of cap and electrodes with Perfektan TB

  - Additional safety information

- Contact addresses

- References

------------------------------------------------------------------------------------------------------------------------

Aim of this document

This document describes the Procedure for measurement of the electroencephalogram (EEG) at the Faculty of Psychology and Neuroscience (FPN). This procedure has been approved by the Ethics Review Committee of Psychology and Neuroscience (ERCPN) and consists of professional standards for safe, hygienic, and valid EEG measurements. It is based on published and generally recognized guidelines (Putnam et al., 1992) and adapted to the local situation. If a researcher explicitly states to adhere to this Procedure in research proposals that are sent to the ERCPN then the ERCPN will generally approve this aspect of the research proposal.

Description of the EEG technique

Electroencephalography (EEG) is a well-established, safe, and relatively comfortable method for clinical and research applications. EEG passively registers electrical activity of the brain. It is typically noninvasive, with the recording sensors placed along the scalp.

Known persistent side- or after-effects

There are no known persistent side- or after-effects. The most often reported and generally short-lived effect is a reddening of the skin due to: a) removal of facial stickers that were placed on skin for holding external sensors; b) preparation of the skin (i.e. rubbing/scrubbing/scratching/abrading) with cotton wool or a cue-tip and/or scrubbing gel (e.g. NuPrep). Skin preparation is usually performed in order to reduce impedance between sensors of interest and a common reference sensor. Note that, lesioning the skin (epidermis) due to overzealous abrading should always be prevented.

Starting an EEG study

In order to be allowed to start an EEG study, a project first needs approval from one of three possible FPN project proposal committees (PPC):

- the EEG PPC

- the NIBS PPC

- the MRI PPC

Specifically:

- Single EEG studies intended to be performed in the dedicated FPN EEG labs or with mobile EEG equipment anywhere else need to go through the EEG PPC.

- Combined EEG/NIBS studies (or single EEG studies intended to be performed in a NIBS lab) need to go through the NIBS PPC

- Combined EEG/MRI studies (or single EEG studies intended to be performed in a MRI lab) need to go through the MRI PPC

The PPCs will check for safety and if formal approval by the Ethics Review Committee Psychology and Neuroscience (ERCPN) has been granted. They also provide feedback on study design, EEG acquisition parameters, and data analysis plan.

You will need to provide a code for your experiment to be used when booking a lab and/ or equipment to keep track of who is using the lab and/or equipment and at which times.

Once formal approval by the EEG, NIBS or MRI PPC is obtained, the PPC will inform the ERCPN. The study will then be registered as “single study” with the ERCPN.

Experimental procedures regarding EEG study

- Participant screening:

In studies involving visual stimulation with flash-like stimuli, having epilepsy or a history of seizures is an exclusion criterion. Also note that some hairstyles (e.g., dreadlocks) make it impossible to achieve a high-quality signal from the scalp.

- Practical training of experimenters

Only users that have received EEG training are allowed to administer EEG measurements. Formal EEG training is offered during scheduled EEG practicals, but can also be given by a knowledgeable supervisor. When you submit an EEG proposal, you will be asked to state who has given (or will give) you training. Note that, besides the procedures described in this document, many other requirements must be met to obtain valid EEG.

Safety guidelines and procedures regarding EEG study

Electrical safety

In order to maintain electrical safety and avoid exposure of participants and researchers to current the following rules need to be taken into account:

- Always use the EEG setup in accordance with the legal standards and internal procedures of FPN / UM.

- Before (first) use of EEG equipment always read the manual of the equipment and make sure to follow the instructions given by the manufacturer and/or Instrumentation Engineering (IE).

- Always connect equipment following the enclosed drawings and /or instructions.

- All equipment used for EEG is CE certified and checked periodic for safety by FPN Instrumentation Engineering (IE). Be aware that any modification to equipment or addition of new equipment can cause (serious) safety issues and/or noise issues.

- Do not connect new or modify existing equipment without consulting IE first. This is especially important when using mobile EEG equipment

Reducing risk for disease transmission

- Wash hands before and after working with a participant.

- Use disposable gloves during skin preparation and cleaning (do not re-use gloves)

- To work in the most hygienic manner, cover the work bench with disposable paper-towel(s). Place cotton swabs, syringes (for the electrode gel), cue-tip/sticks for skin preparation, etc. only on this sheet (do not re-use the sheet).

- Only use new and clean items to prepare the skin (do not re-use items)

- Dispose items that have been in contact with skin immediately after use

- Blunt needles that are used to insert gel into the sensors should be disposed in a puncture resistant container

- Sensors and cap should be cleaned gently with a soft (tooth-) brush and (distilled) water

- Disinfection of sensors and caps after every use is required when skin/scalp has been rubbed/scrubbed/scratched/abraded.

- Perfektan TB is recommended for disinfecting (see specific instructions below) but other types of disinfectants can be used as long as they are as effective against blood borne viruses and bacteria as 15 minutes of submerging in a 3% Perfektan TB solution is.

- Disinfection is not required (but still recommended) when EEG measurements are performed without the skin/scalp being rubbed/scrubbed/scratched/abraded. In this case, cleaning with ivory soap is sufficient. Typically, not rubbing/scrubbing/scratching/abrading the skin only leads to satisfactory data quality if active electrodes are used, since they permit higher electrode-impedances. Note that when applying gel is not sufficient to reach the desired electrode-skin contact and impedance-reducing actions were performed, disinfection of caps and electrodes (active and passive) is required as skin damage cannot be ruled out.

Chemical safety

In order to maintain chemical safety and avoid exposure of participants and researchers to hazardous materials only use chemicals following the legal standards and internal procedures of UM / FPN.

Specific instructions for disinfection of cap and electrodes with Perfektan TB

Instructions for use in the dedicated FPN EEG labs (UNS40, 4th floor):

Personal protection

- PerfektanTB is toxic and irritating: ALWAYS wear protective gloves when handling Perfektan TB

- Avoid spraying and aerosol formation

- Avoid contact with skin and eyes: ALWAYS wear safety goggles

Clean caps and electrodes with water

    - Clean caps with water, use a toothbrush to remove gel from the cap

    - Clean electrodes with water and softly remove gel with toothbrush

Preparing Perfektan TB

- The scale of the solution is: 30 ml of Perfektan TB and 970 ml of water (3%). Use glass measuring cans to prepare the solution.

- Always work in the fume cupboard when handling Perfektan TB

- The solution should be freshly prepared daily*, but not more often, so it can be re-used. Add a yellow note with the date of the preparation of solution. If you come in and see that the solution is from the day before, please put it in the designated can (see below) and make a new solution and add the date of preparation

- The glass measuring can with the solution in it has to stay in the fume cupboard at all times

- Ensure that the caps and electrodes are completely immersed and the solution has contact with all surfaces. Avoid cables and connectors getting wet (!)

Contact time

- Remove the caps and electrodes after 15 minutes (NOT LONGER!!!) and then rinse both thoroughly with water in the cabinet (use the green handle). Always rinse the measurement cups with (distilled) water after each use.

Chemical waste Category 3

- After use, the solution has to be collected in 10 liter cans labeled with “Category 3” and “Perfektan TB solution”. Do not throw this solution in the sink!

- A 10 liter can with solution in it always has to be stored in the cabinet below the fume cupboard. Empty cans can also be found in that cabinet

- All Perfektan TB bottles (full, in use or empty) always have to be stored in the yellow safety-storage cabinet at all times

Instructions for use anywhere outside of the dedicated FPN EEG labs:

Personal protection

- PerfektanTB is toxic and irritating: ALWAYS wear protective gloves when handling Perfektan TB

- Avoid spraying and aerosol formation

- Avoid contact with skin and eyes: ALWAYS wear safety goggles

Clean caps and electrodes with water

    - Clean caps with water, use a toothbrush to remove gel from the cap

    - Clean electrodes with water and softly remove gel with toothbrush

Preparation of Perfektan TB

- Prepare the solution in the fume cabinet: 30 ml of Perfektan TB and 970 ml of water (3%). Use glass measuring cans to prepare the solution.

- Transport Perfektan TB solution in a closed container to measurement location

- The used solution should be freshly prepared daily*, but not more often, so it can be re-used

- Place the used cap and sensors into a separate container and fill with pre-prepared Perfektan TB solution

- Ensure that the caps and electrodes are completely immersed and the solution has contact with all surfaces. Avoid cables and connectors getting wet (!)

Contact time

- Remove the caps and sensors after 15 minutes (NOT LONGER!!!) and then rinse both thoroughly with water. Always rinse the measurement cups with (distilled) water after each use.

Chemical waste Category 3

- After use, the solution has to be collected in 5 liter cans labeled with “Category 3” and “Perfektan TB solution”. Do not throw this solution in the sink!

- Take used Sekusept solution back to UM in a closed container and deport as chemical waste.

OR when being off campus and no fume cupboard is available:

Preparation of Perfektan TB

- Prepare diluted Perfektan TB solution in a well-ventilated room on location (30 ml of Perfektan TB and 970 ml of water, 3%)

- Place the used cap and sensors into a separate container and fill with pre-prepared Perfektan TB solution

- The used solution should be freshly prepared daily*, but not more often, so it can be re-used

- Ensure that the caps and electrodes are completely immersed and the solution has contact with all surfaces. Avoid cables and connectors getting wet (!)

Contact time

- Remove the caps and sensors after 15 minutes (NOT LONGER!!!) and then rinse both thoroughly with water. Always rinse the measurement cups with (distilled) water after each use.

Chemical waste Category 3

- After use, the solution has to be collected in 5 liter cans labeled with “Category 3” and “Perfektan TB solution”. Do not throw this solution in the sink!

- Take used Perfektan TB solution back to UM in a closed container and deport as chemical waste.

First aid measures

General information: 	Remove contaminated soaked clothing immediately. If you feel unwell, seek medical advice.

After skin contact: 		Wash off immediately with soap and plenty of water. Consult a doctor if skin irritation persists.

After eye contact: 		Rinse immediately with plenty of water, also under the eyelids, for at least 15 minutes. Remove contact lenses, if present and easy to do. Consult (eye) doctor immediately.

After ingestion: 		Do not provoke vomiting. Consult physician. Attention in case of vomiting – acute danger of suffocating, produced by foaming ingredients. Rinse mouth. Make drink some glasses of water. The decision whether to provoke vomiting is to be taken by a physician.

After inhalation: 		Move to fresh air in case of accidental inhalation of vapours or decomposition products. In the event of symptoms refer for medical treatment.

Most important symptoms and effects, both acute and delayed

- Causes skin irritation.

- Causes serious eye damage

- Watch out. Beware, hazard of foam aspiration.

Safety information

- H302	Harmful if swallowed

- H314	Causes severe skin burns and eye damage

- H315	Causes skin irritation

- H318 	Causes serious eye damage

- H319	Causes serious eye irritation

- H335	May cause respiratory irritation

- H373	May cause damage to organs through prolonged or repeated exposure

- H400	Very toxic to aquatic life

- H410	Very toxic to aquatic life with long lasting effects

- H412	Harmful to aquatic life with long lasting effects

Ingredients Perfektan TB (aldehyde-free)

100 g contains: 5.6 g N, N-didecyl-N-methyl-poly(oxyethyl)ammonium propionate;

3,75 g cocospropylenediamine-1,5-bis-guanidinium diacetate. Contains fragrances:

Limonene, 5–15% non-ionic surfactants, disinfectant.

Additional Safety Information

- Please keep the dedicated EEG labs, as well as the general lab area and closets, neat and clean

- Note that according to UM Health and Safety policy it is not allowed to bring food or beverages into the lab areas

Addresses

Lab coordinator for EEG-1 to EEG-4 and mobile EEG labs

- Franc Donkers (3885260), franc.donkers@maastrichtuniversity.nl, OXF55-2.009

Agenda for EEG-1 to EEG-4 and mobile lab EEG labs

- https://resourcebooker.maastrichtuniversity.nl/

Server location for useful EEG lab information

- \\unimaas.nl\Research\FPN\EEG-Labinfo\

Instrumentation

- Head: Jacco Ronner, j.ronner@maastrichtuniversity.nl, UNS40-3.748

- Johan Gielissen, j.gielissen@maastrichtuniversity.nl, UNS40- 3.777

- Luc Offermans, l.offermans@maastrichtuniversity.nl, UNS40- 3.777

- Tim Uiterwijk. tim.uiterwijk@maastrichtuniversity.nl UNS40-3.777

Ethical approval

- Ethical approval is required for all studies conducted under the responsibility of FPN staff: See: https://www.maastrichtuniversity.nl/about-um/faculties/psychology-and-neuroscience/facilities/ethical-review-committee-psychology-and

References

Putnam, L.E., Johnson, R., & Roth W. (1992). Guidelines for reducing the risk of disease transmission in the psychophysiology laboratory. Psychophysiology, 29(2), 127-141.